**Turbo Mixer Troubleshooting**

When it comes to the Turbo Mixer and the issue of running too fast or too slow, the issue is almost always attributed to the **Potentiometer**;this can be found at the bottom of the unit and can be accessed with a small Philips head screwdriver.

1. Ensure that the unit is powered on and the speed knob is turned all the way down
2. Use a screwdriver to access the potentiometer and turn right **VERY SLOWLY,** as it does not take much to get the motor going
3. Make adjustments until you have the Plus/Minus .5% of 3000 RPMs
4. Test unit.

**If you have any questions, please contact our Product Specialist Team at 800-726-7345 Option 2 or Service Department at 800-726-7345 EXT 4348.**