Purpose

This work instruction describes the process for replacing the MXU Solenoid.

These repair instructions are valid for revision C MXU Digital Centrifuges. For revisions older than C, please consult Product Development.

Equipment/Software

| ITEM # | PART # | DESCRIPTION | QTY. |
| --- | --- | --- | --- |
| 1 | 100079 | Solenoid; 12V | 1 |

Phillips head screwdriver

#7 nut driver or #7 socket with socket wrench

Instructions

Confirm that all necessary components are present to complete the conversion.

Open the lid and remove the rotor and tube shields from the unit. Figure 1

Figure 1

Turn the unit over onto its lid and remove the base plate. Figure 2

- NOTE: Avoid overextending connected wires when removing the base plate.

Figure 2

Disconnect the Motor (J3), the Speed IR Sensor (J5), and the Solenoid from the Motor Drive Board. Figure 3

Solenoid

J3

J5

Figure 3

Using the #7 nut driver, locate and remove the two nuts (100922) holding the Solenoid IR Sensor Bracket (101110) in place. Figure 4

101110 – Solenoid IR Sensor Bracket

Figure 4

Using a Phillips head screwdriver, remove and replace the Solenoid (100079) from the Solenoid IR Sensor Bracket. Figure 5

Figure 5

Re-attach the Solenoid IR Sensor Bracket to the body. Figure 6

Figure 6

Re-connect the Motor (J3), Speed IR Sensor (J5), and the Solenoid to the Motor Drive Board. Figure 7

Solenoid

J3

J5

Figure 7

Re-attach the Base Plate Assembly to the Body. Figure 6

Figure 6

Replace Rotor, close the lid, and verify the unit powers on and functions under normal laboratory use.