Purpose

This work instruction describes the process for replacing the MXU Power Switch (100413).

These repair instructions are valid for revision C MXU Digital Centrifuges. For revisions older than C, please consult Product Development.

Equipment/Software

| ITEM # | PART # | DESCRIPTION | QTY. |
| --- | --- | --- | --- |
| 1 | 100413 | Switch; Power Input | 1 |

Phillips Head Screwdriver

Instructions

Confirm that all necessary components are present to complete the conversion.

Open the lid and remove the rotor and tube shields from the unit. Figure 1

Figure 1

Turn the unit over onto its lid and remove the base plate. Figure 2

- NOTE: The motor wiring harnesses will be connected. Disconnect them or set the base close enough to the side to not overextend the wires.

Figure 2

Figure 3

3.4 Locate Power Switch on the front face of the unit. Figure 3

Disconnect the 2 cables from the back of the Power Switch (100300-B & 100325-B). figure 4

100325-B

100300-B

Figure 4

While pressing inwards on the top and bottom tabs of the Power Switch, push the switch out through the front face of the body. Figure 5

Figure 5

Insert new Power Switch, ensuring the “|” symbol faces up (towards the lid), and that the two tabs snap in place.

Reconnect the 2 cables as shown. Figure 6

100325-B

100300-B

Figure 6

Re-attach the base plate. Figure 7

Figure 7

Replace Rotor, close the lid, and verify the unit powers on and functions under normal laboratory use. Figure 8

Figure 8