Purpose

This work instruction describes the process for replacing the Lid Gasket of MX12 centrifuge units.

Equipment/Software

| ITEM # | PART # | DESCRIPTION | QTY. |
| --- | --- | --- | --- |
| 1 | 100993 | Rubber Gasket | 1 |

Instructions

Confirm that all necessary components are present to complete the conversion.

Replacement Lid Gasket

- Reference Figure 1

Power on your MX12 unit and open the lid. Then turn off the power to your unit.

Remove the old Lid Gasket (100993) from your unit and discard it.

Carefully place your new replacement Lid Gasket (100993) onto the rim of the bowl opening, referencing Figure 2.

- Reference Figure 2

Firmly press Lid Gasket (100993) over the rim of the bowl opening, ensure that the cushioned side is facing upwards, referencing Figure 3.

- Reference Figure 3

This completes the repair.