Purpose

This work instruction describes the process of replacing the MX5 lid latch.

Equipment/Software

| ITEM # | PART # | DESCRIPTION | QTY. |
| --- | --- | --- | --- |
| 1 | 100912-D | MX5 Lid Latch | 1 |

Tools Needed: 2.5mm Allen Wrench/Driver, Loctite Blue.

Instructions

Confirm that all necessary components are present to complete the conversion.

Ensure the unit is powered off and unplugged.

Open the MX5 Lid. Figure 1

Figure 1

Using the 2.5 Allen Wrench/driver, remove the 4 M4 screws (101071). Figure 2

Figure 2

M4 Screws (101071)

MX5 Lid Latch (100912-D)

Apply a small amount of Loctite Blue to the 4 M4 screws (101071) and replace the MX5 Lid Latch (100912-D). Figure 3

Figure 3

Close the lid.

Plug the MX5 in and power it on. Ensure unit functions under normal laboratory use.

Repair is complete.