1. Purpose
   1. This work instruction describes the process of replacing the Fuse on the MX12 centrifuges.
2. Equipment/Software

|  |  |  |  |
| --- | --- | --- | --- |
| ITEM # | PART # | DESCRIPTION | QTY. |
| 1 | 100185 | Fuse | 1 |

* 1. TOOLS NEEDED – N/A

1. Instructions
   1. Confirm that all necessary components are present to complete the conversion.
   2. Turn off your MX12 unit and disconnect from power.
   3. Unplug the old Fuse (100185) and discard it. Figure 1

A black electrical outlet with a switch

Description automatically generated­­­

100185

100053

Reference Figure 1

* 1. Plug in the new Fuse (100185).
  2. Close the lid. Run the MX12 unit at 12,000 RPM for 3 minutes to verify successful replacement of the new Fuse (100185).
  3. This completes the repair.