Purpose

This work instruction describes the process for replacing the MXU Speed Wheel.

These repair instructions are valid for revision C MXU Digital Centrifuges. For revisions older than C, please consult Product Development.

Equipment/Software

| ITEM # | PART # | DESCRIPTION | QTY. |
| --- | --- | --- | --- |
| 1 | 101103 | Speed Sensor Wheel | 1 |
| 2 | 101104 | Speed Sensor Hub | 2 |
| 3 | 101105 | Speed Sensor Hub Stop | 1 |
| 4 | 100787 | Blind Rivet, SS | 2 |

Phillips Head Screwdriver

Instructions

Confirm that all necessary components are present to complete the conversion.

Open the lid and remove the rotor and tube shields from the unit. Figure 1

Figure 1

Turn the unit over onto its lid and remove the base plate. Figure 2

- NOTE: Avoid overextending connected wires when removing the base plate.

Figure 2

Disconnect the Motor (J3) and the Speed IR Sensor (J6) from the Motor Drive Board. Figure 3

J3

J6

Figure 3

Remove the Motor Assembly from the Base Plate. Set the Base Plate aside. Figure 4

Motor Assembly

Base Plate

Figure 4

Remove the Speed Wheel Assembly from the Motor Assembly. Figure 5

Speed Wheel Assembly

Figure 5

Speed Wheel Assembly

- Align the Speed Wheel (101103), Speed Wheel Hubs (101104), and Speed Wheel Hub Stop (101105) as shown in Figure 6.

- Ensure that all pieces lay flat together and no burrs are present.

- Rivet the pieces together with Blind Rivets (100787)

| ITEM # | PART # | DESCRIPTION | QTY. |
| --- | --- | --- | --- |
| 1 | 101103 | Speed Sensor Wheel | 1 |
| 2 | 101104 | Speed Sensor Hub | 2 |
| 3 | 101105 | Speed Sensor Hub Stop | 1 |
| 4 | 100787 | Blind Rivet, SS | 2 |

Figure 6

Re-attach the new Speed Wheel Assembly to Motor Assembly. Figure 7

Figure 7

Re-attach Motor Assembly to Base Plate. Figure 8

Figure 8

Motor Assembly

Base Plate

Re-connect the Motor and Speed IR Sensor to the Motor Drive Board. Figure 9

J3

J6

Figure 9

Re-attach the Base Plate to the Body. Power the unit on and verify it functions under normal laboratory use. Figure 10

Base Plate

Body

Figure 10